## Slide 1: Optics Transmission Update

## Slide 2: Importance of Exposure
Niko estimated the transmission through the rings to be 1.5% and 3%
This requires ~100x “dynamic range”
Due to the paper structure*, given a single exposure time, there is only a range of ~3x
*and maybe something else

## Slide 3: Exposure Test
Also had to include power meter uncertainty for good fits (0.015 W + 0.7 %)
Tested Exposure Times from 1 sec to 3 min
Initial linear fits statistically failed, but carefully controlling the repositioning of the camera correct this.

## Slide 4: Results
Tested Exposure Times from 1 sec to 3 min
All of them consistent to at least 3%
3 sec (last weeks data)
1.5 sec
6 sec
180 sec
The data set with better controlled positioning gave ~0.5% consistency

## Slide 5: Summary
Power response and Exposure times are linear to at least ~1% over a wide span (0.5 to 100 W and 1.5 to 180 sec) for a fixed region
How do different regions compare?

## Slide 6: Positioning Test

## Slide 7: Off Center Blur

## Slide 8: Yikes
Factor of 1.5 different (~132)

## Slide 9: Next Step: Beam Divergence